## Sheet1
BONE DENSITOMETER, FAN BEAM AND LINEAR SCANNING
ITEM | TECHNICAL SPECIFICATIONS
1 | Modality Details
1.1 | Functionality:
1.2 | The system must be a Linear Scanning device.
1.3 | 2D Fan beam Technology
1.4 | Motorized Table Top and Arm.
1.5 | Laser light for patient positioning
1.6 | Maximum patient weight of 200kg or above.
1.7 | Table lowest height range between 60cm - 72cm.
1.8 | A complete scan / exposure time for AP spine & femur and hip study should be completed within 40 seconds.
1.9 | Multi-element, high resolution, digital detector array. The details of the detector must be stated.
1.10. | The maximum scanning area must be a maximum range of between 60-75 cm in width.
1.11 | The maximum scanning area must be a maximum range of between of 190-210 cm in length.
1.12 | Dual Energy X-Ray Absorptiometry (DEXA) Dual pulse energy.
1.13 | The system must be able to generate a maximum kVp of 100 or above.
1.14 | Both Automatic and Manual selection of Region of Interest must be possible.
1.15 | Automatic and continouous calibration. Operator calibration not required.
1.16 | Automatic quality control program, including required phantoms.
1.17 | In addition to its normal operation, the system must be suited to the examination of paediatric patients.
1.18 | OPERATOR CONSOLE
1.18.1 | Must come with a PC with the following requirements:
1.18.1.1 | Windows 10 or newer operating system.
1.18.1.2 | i7 or better processor.
1.18.1.3 | 32 GB RAM
1.18.1.4 | 4 TB hard drive.
1.18.1.5 | 17 Inch or larger 3 MP Flat Panel Touch Monitor.
1.18.1.6 | Keyboard and mouse
1.18.1.7 | CD/DVD R/W drive enabled (it should be license activated and ready for CD/DVD writing)
1.18.1.8 | Network ready
1.18.1.9 | DICOM compliant and enabled (HIS, RIS,PACS).
2 | Software tools and functions
2.1 | The following Software Packages must be included in the offer:
2.1.1 | Bone Mineral Density (BMD), including hip and Lateral Spine BMD.
2.1.2 | Bone Mineral Content (BMC)
2.1.3 | Body Composition Analysis
2.1.4 | BMI Calculator
2.1.5 | Scan / analyze protocols for whole body, AP and lateral spine, proximal femur, dual hip, forearm, scoliotic spine for paediatric and adult patients.
2.1.6 | Digital vertebra assessment (IVA).
2.1.7 | The system must have reporting packages.
2.1.8 | The reporting packages must be DICOM compatible.
2.1.9 | Customizable reports.
2.1.10 | The system must support multiple reports per patients, including comparison against earlier studies.
2.1.11 | The reporting software must be compliant with the International Society for Clinical Densitometry (ISCD).
2.1.12 | Integrated FRAX report to be included (fracture risk assessment).
2.1.13 | The system must include a compatible Colour Printer.
4 | Operating Requirements
4.1 | Temperature: 15 - 32°Celsius
4.2 | Humidity: 20 - 80%
4.3 | Power: 220 - 240 V, 50 Hertz
5 | IT Integration
5.1 | The system must have the following DICOM compatibility:
5.1.1 | DICOM print class
5.1.2 | DICOM storage class
5.1.3 | DICOM query / retrieve
5.1.4 | DICOM work list
6 | ACCESSORIES
6.1 | forearm positioner with palm mouse positioner
6.2 | head positioner
6.3 | hip positioner fixture with a velco strap
6.4 | spine phantom
6.5 | step phantom for vertebral assessment
6.6 | positioning sponges
7 | POWER SUPPLY
7.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
7.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
7.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
7.4 | A 3 phase in / out power supply must be offered. State details.
7.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
7.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
7.7 | The lifetime of UPS must be at least 15 years.
7.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
7.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
7.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
7.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
7.12 | Maintenance free batteries must be offered.
8 | RADIATION PROTECTION
8.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
8.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
8.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
8.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
8.5 | A radiation report must be automatically generated and transferable via DICOM.
8.6 | Should be FDA or CE approved product
8.7 | Electrical safety conforms to standards for electrical safety IEC-60601
8.8 | The unit must comply to IEC-61267
8.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
8.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
8.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
8.12 | Lead Glass window to be installed as per room requirement
8.13 | Safety aspects of Radiation dosage leakage should be spelt out
8.14 | DB boards and electrical filters to be installed
8.15 | QA TOOLS
8.16 | Display Exposure index values
8.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
8.18 | Patient Conventional face shield to be included.
8.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
8.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
8.21 | Interlocks to stop system movement when compression is applied must be provided.
8.22 | Interlock for premature release of the X-ray switch must be available
8.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
8.24 | Phantom for calibration should be provided.
9 | AIR CONDITIONER
9.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
9.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
9.3 | The air conditioner must also be included in the maintenance plan.
9.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
9.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
9.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
9.7 | Air-conditioning must be included in Maintenance plan
8 | Optional Items
8.1 | Orthopaedics software for prosthesis.
8.2 | Orthopaedic software
8.3 | Single energy femur scan software- to detect atypical femur fractures