## Sheet1
Diagnostic OAE with standard frequency
Description
Device used for diagnostic assessment of oto-acoustic emissions within the cochlear region.
Technical Specification
Advanced clinical oto-acoustic emission system, able to record and archive completed tests
Must be able to perform TEOAEs and DPOAEs
DPOAE frequencies: 1.5- 6 kHz, Intensities: 40-70 dB SPL SNR: 3-10 dB, Averaging Time: 0.5,1,2,4 sec
TEOAE frequencies: 0.7-4 kHz Intensity: 83 dB pe SPL SNR: 3-10 dB, Averaging Time: 8, 16, 32, 64 sec
Must include the following:
bilateral probes
probe cable clips
nubs and probe pouch
USB cable to computer connection
Must be able to indicate noise flow measurements
Laptop requirements are a minimum of 5GB free disk space, minimum 1GB of RAM, Latest Intel version (for the option of connecting to a laptop or PC for printing purposes)
Accessories
bilateral probes
probe cable clips
nubs and probe pouch
USB cable to computer connection
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list