## Sheet1
Combination Electrotherapy and Ultrasound Unit
Description
A multimodal electrotherapy unit used for pain relief, muscle stimulation and to promote healing
Technical Specification
The unit shall offer ultrasound and electrotherapy combination therapy
The combination electrotherapy unit shall include 3 modalities: electrotherapy, ultrasound and magneto therapy
Electrotherapy modality shall provide for at least the following current types: Interferential, TENS, Russian stimulation, diadynamic and galvanic current
The unit shall have at least two independent electrotherapy channels
Intensity: 0-3/cm2 (ultrasound); 0-50mA (IF);0-100MA (LF)
Ultrasound output: 1&3 MHz frequencies to be available
Ultrasound head: 5cm2
Pulse Width: 30-400 microseconds
Pulse rate-carrier frequency: 2KHz-10KHz for IFC; 2.5KHz for Russian; 0Hz-250Hz for TENS and EMS
Pulse rate-beat frequency: 1Hz-200Hz for IFC; 20Hz-100Hz for Russian; 0Hz-250Hz for TENS and EMS
Working time: 60minutes
The unit shall be equipped with body part navigation feature
The unit shall be equipped with a patient database
The unit shall have the option of 2 or 4 pole interference
The unit shall have a built in vacuum unit
The unit shall be battery operated for ease of portability
All accessories necessary to perform all the functionalities of the unit shall be included in the offer
Accessories:
Laser clusters
Ultrasound head 5cm
Ultrasound head 1cm
Infrared laser probes
Vacuum electrodes
Hands free sono applicator
Rubber electrodes
Adhesive electrodes
Russian stimulator
5cm2 ultrasound head
1cm2 ultrasound head
Adapter 100-240V; 47-64Hz
Power cable
User manual
Self adhesive electrodes in various sizes (50x100mm; 50x50mm)
Carbon electrodes in various sizes (70x110mm; 70x100mm)
Envelope sponges to fit carbon electrodes
Skin leads wires
Straps 75x1200mm and 75x600mm
Combo lead wires
Optional extras
Vacuum unit compatible with the supplied unit
Optional Accessories:
2 tier trolley
The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS
The unit shall be FDS or CE approved (Attach certification)
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list