## Sheet1
PET CT (64 slices)
Item | Technical Specifications
1 | PHYSICAL FEATURES
1.1 | A full ring PET system plus integrated CT scanner offering 64 acquired CT slices per rotation including a patient bed and single combined PET/ CT acquisition console.
2 | Gantry
2.1 | PET Gantry
2.1.1 | For minimized count rate losses and high detection efficiency the detector crystal should consist of a lutetium based scintillator.
2.1.2 | The proposed PET scanner must be capable of acquiring at least 81 transverse cross-sectional slices simultaneously without undergoing any axial motion. (TrueV, 109 slices)
2.1.3 | The PET detector elements must be isotropic and have axial and transaxial dimensions less than or equal to 4 mm
2.1.4 | The geometric axial field of view (FOV) as measured from the outer edges of the crystals in the axial direction must be >16.2 cm. (TrueV, 22.1)
2.1.5 | The Gantry must have a 78 cm or greater aperture throughout the entire length of the tunnel for patient comfort, ease of positioning and RTP support.
2.1.6 | Detector design must employ a modular design for ease of service and field upgradability to a different size field of view.
2.1.7 | The (center-to-center) separation between slices acquired simultaneously without any axial motion must not exceed 2.0 mm.
2.1.8 | The scanner must have a Coincidence Timing Window no greater than 4.1 nsec.
2.1.9 | The scanner must have an Energy Window with lower level discriminator greater than or equal to 435 keV.
2.1.10 | The number of pixels (detector elements) per PMT must be less than 45 pixels/PMT.
2.1.11 | The system should contain 24,336 or more individual crystal pixels (32,448 with TrueV).
2.1.12 | The scanner must not use gantry detector wobble or axial shift of the patient to obtain the specified transaxial and axial specifications.
2.2 | CT Gantry
2.2.1 | The CT should support spiral scanning with a maximum spiral acquisition of 100 seconds or greater.
2.2.2 | The transverse field-of-view for CT used during PET attenuation correction should extend all the way to the edges of the gantry and offer uniformity to within 50 HU when imaging greater than 50 cm FOV.
2.2.3 | CT tube must cool at a rate of 5 MHU/min or greater
2.2.4 | The CT should provide at least 64 acquired slices
2.2.5 | The CT should have a 360 degree rotation speed of at least 0.30 seconds.
2.2.6 | The CT must have spiral (pitch) that is freely variable without a change in image quality.
2.2.7 | CT image quality must provide <0.4 mm resolution scanning for all clinical scans, and < 0.25 mm resolution in hi resolution mode.
2.2.8 | The CT system must be capable of 30 LP/cm
2.2.9 | The CT subsystem will have a dynamic collimator dose shield reduction technique to reduce pre and post scan dose to the patient
2.2.10 | The CT must have on the fly, real time dose modulation without additional radiogram procedures.
2.2.11 | The CT must be capable to scan at 70 kV for reduced dose to pediatrics
2.2.12 | The CT generator power must be at least 100 kW.
2.2.13 | The gantry thickness should not exceed 136 cm.
2.2.14 | The PET/CT gantry must require less than 7000 BTU/hr for HVAC room cooling.
2.2.15 | The gantry should have a closed design without vents/fans to reduce ambient noise and temperature restrictions for the scan room.
3 | Patient Bed
3.1 | The horizontal motion of the patient bed must be electrically motorized and computer-controlled with an independent operator control option as well. Operator controls accessible from both sides of the patient must be provided for both horizontal and vertical control.
3.2 | The patient bed must support the use of continuous bed motion during acquisitions to eliminate PET bed stops.
3.3 | A digital readout of the horizontal and vertical position of the bed must exist and must be located near the operator controls for the bed to provide ease in positioning.
3.4 | The patient bed must provide a constant deflection throughout the scanner axial field-of-view. The bed should support a maximum patient weight of 230 kg or greater at full scan coverage.
3.5 | The lowest vertical position should be 65 cm or less from the floor to the top of the side edge of the patient pallet.
3.6 | The horizontal travel of the bed should enable a co-scan axial field-of-view of 190 cm or greater.
3.7 | A method of preventing accidental vertical motion by an extraneous activation of the operator's controls must be provided to prevent the hazard of patient injury.
3.8 | A method of stopping any computer controlled or accidental motion must be provided and must be located at all operator positions of the system.
3.9 | A mechanism for manually removing the patient bed from the gantry bypassing the electric motor control must be supplied for emergency situations and described fully in the proposal.
3.10. | The patient bed should support the use of a flat RTP pallet (width > 50 cm)
3.11 | The bed should be usable for Radiotherapy threatment planning
4 | SOFTWARE
4.1 | Data Acquisition:
4.1.1 | The software must exist to acquire data in 3D (e.g., as a volume data set) modes. Reconstruction by the FORE and OSEM iterative algorithms must be included.
4.1.2 | The acquisition program should support pre-programmed scan protocols with acquisition and reconstruction parameters and patient information with simple, dynamic editing of parameters. These parameters would include all information necessary to acquire data on the PET scanner (e.g., scan duration, patient information, bed motion), as well as information necessary for reconstruction, which would be carried through to the reconstruction step.
4.1.3 | Continuous bed motion during a whole body PET scan is required.
4.1.4 | Whole body with multi-bed and/or continuous bed motion acquisition protocols should be performed such that operator interaction is not required to complete the scan once it has been initiated.
4.1.5 | As a part of static acquisition, a total count rate display should be updated at least every 5 sec.
4.1.6 | The system must include a single, integrated console, which controls acquisition for both the CT and PET.
4.1.7 | The software must support DICOM (PET and CT) export, import, and well as filming and archival.
4.1.8 | Online data storage capacity should be a minimum of 1 TB
4.1.9 | Metal artifact reduction must be included for the CT and the new images with reduced artifact, must be used for attenuation correction
4.1.10 | A raw data iterative CT reconstruction technique must be included with the CT system
4.1.11 | The system must have the ability to cover 9 cm of perfusion data during Perfusion CT studies
4.1.12 | The system must support dual energy CT.
5 | Reconstruction and Reformatting
5.1 | Static image reconstruction should commence automatically upon completion of a bed position or acquiring enough information during CBM to begin parallel with the acquisition, using the reconstruction parameters defined at acquisition.
5.2 | If the system cannot keep pace with the acquisition, there should be a method for selecting a specific scan to be reconstructed immediately.
5.3 | In that case, other scans should be reconstructed immediately after the requested reconstruction has finished.
5.4 | The methods for randoms estimation and subtraction must be by measuring randoms rather than calculations from measured singles.
5.5 | The system must include algorithms for utilizing the CT data to generate attenuation correction for the PET data.
5.6 | A PET frame (200x200) should be reconstructed in < 30 sec/bed with all data corrections, 3D AW-OSEM
5.7 | 512*512 CT images should be reconstructed at a rate of at least 40 images per second
5.8 | The user should have the capability to specify the transverse FOV for reconstruction.
5.9 | The reconstruction program should support reconstruction in image sizes of at least 360x360 for PET and 512x512 for CT.
5.10. | If the pixel and reconstruction matrix sizes are selected such that the reconstructed area is smaller than the transaxial FOV, it should be possible to select where in the transaxial FOV the reconstruction is centered.
5.11 | A scatter correction technique that is spatially variant and adjusts for patient geometry must be included. Scatter correction must be provided based on scan of the actual patient whose scan is being corrected and processed automatically in an iterative loop
5.12 | A fully automated, one-button protocol for whole body oncology studies should be provided, having the following features:
5.12.1 | permits multi-bed or CBM scans to proceed in an efficient way without operator intervention;
5.12.2 | processing software which proceeds automatically without operator intervention (processes include segmented attenuation correction, iterative reconstruction, scatter correction, normalization, smoothing, filtering, etc.)
5.12.3 | display of cine whole body projection views as well as coronal, transverse, and sagittal volume slices including multiple reconstruction jobs, auto archiving and networking.
5.13 | The system must be capable of performing a Full 8 gate cardiac perfusion/MBF acquisition with complete reconstruction of all the data necessary for the rest and stress portion of the study in less than 23 minutes from start to finish, barring no pharmaceutical stress issues.
5.14 | Measured Point Spread Functions of the system must be utilized in the iterative reconstruction algorithm.
5.15 | Respiratory gating must be included which can automatically find and automatically reconstruct a motion free image from the gated PET data.
5.16 | Cardiac gating must be included which can correct for respiratory motion using a deviceless method
5.17 | Gating inputs must be located on the patient bed to reduce hanging wires for safety.
5.18 | Gating device must be interfaced with PET/CT system eliminating the need to access the gating laptop for the UI in order to acquire data.
5.19 | A whole body scan must be fully capable of integrating respiratory gated acquisition/data into both the acquisition and display
5.20. | PET and CT respiratory gated data can be automatically reconstructed in a phased matched manner.
6 | Display and Analysis
6.1 | It is required that the acquisition, reconstruction and image display and analysis software to operate in Windows 10 or greater. A description of the software environment and flexibility should be given in the proposal.
6.2 | The system must be able to display CTDIw and DLP
6.3 | Region of interest (ROI) capability for circular, rectangular, and irregular ROI's must be provided.
6.4 | A patient database with such entries as patient name, patient weight, injected activity and time, identification number, protocol (disease), date of scan, data filename is required and should be discussed in the proposal.
6.5 | Utility programs for access to database information should also be discussed.
6.6 | The Patient Database should support DICOM worklist functionality.
6.7 | A visualization application that allows the comparison of 3 fused datasets (e.g. pre- and post- therapy PET/CTs) on the same screen with all gating and quantification capabilities.
7 | Quality Control
7.1 | PET QC can be set up and run automatically overnight (optional).
7.2 | A procedure to calibrate the scanner to a well counter must be provided with the system and described in the proposal.
7.3 | The QC phantoms must be part of the unit (optional).
7.4 | Daily quality control procedures are well defined and automated, and all necessary sources are included with the system
7.5 | An apparatus for reproducible positioning is required and should be described in the proposal.
8 | PET Performance Specifications
8.1 | Transverse Resolution
8.2 | Transaxial resolution (in-plane resolution) specification for 3D mode must be equal to or better than 4.5 mm FWHM at 1 cm from the central axis of the gantry.
8.3 | Axial Resolution
8.4 | Axial resolution specification for 3D mode must be equal to or better than 4.7 mm FWHM at 1 cm from the central axis of the gantry.
8.5 | Sensitivity
8.6 | System sensitivity specification must exceed 10.2 cps/kBq at 435 keV LLD.
8.7 | Count Rate
8.8 | The count rate peak NECR performance must achieve at least 180 kcps
8.9 | System must offer a 3D PSF PET reconstruction technique
8.10. | Time of flight capability must be available with better than 550 ps timing resolution at all clinical count rates
9 | POWER SUPPLY
9.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
9.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
9.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
9.4 | A 3 phase in / out power supply must be offered. State details.
9.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
9.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
9.7 | The lifetime of UPS must be at least 15 years.
9.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
9.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
9.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
9.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
9.12 | Maintenance free batteries must be offered.
10 | Miscellaneous
10.1 | Software for archival and retrieval of scan and image data must be provided and described in the proposal.
10.2 | The system shall support the use of USB drives for fast easy storage