## Sheet1
Screening tympanometer (226Hz and 1000Hz)
Description
Portable screener for tympanometry and ipsilateral acoustic reflex testing
Technical Specification
Tympanometry measurement capability of 226Hz.
Must test tympanometry across range of +200 -to -400daPa with accuracy of ±10daPa
Ipsi-lateral testing range from 500,1000,2000 and 4000Hz
Must include probe/ear tips
Must be supplied with protective carry case
Must have A/C power supply
Must have printing capabilities and supplied with two rolls of paper
Must have a memory for at least 12 tests, with probe assembly
Must be supplied with a box of ear tips of at least six different sizes, with at least three tips of each size
Accessories:
probe/ear tips of at least six different sizes, with at least three tips of each size
protective carry case
two rolls of paper
memory for at least 12 tests
probe assembly
box of ear tips
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list